**Best Practices for Selecting a Predicate Device to Support a Premarket Notification [510(k)] Submission**

**Draft Guidance for Industry and Food and Drug Administration Staff**

***DRAFT GUIDANCE***

**This draft guidance document is being distributed for comment purposes only.**

**Document issued on September 7, 2023.**

You should submit comments and suggestions regarding this draft document within 90 days of publication in the *Federal Register* of the notice announcing the availability of the draft guidance. Submit electronic comments to [https://www.regulations.gov](https://www.regulations.gov/). Submit written comments to the Dockets Management Staff, Food and Drug Administration, 5630 Fishers Lane, Room 1061, (HFA-305), Rockville, MD 20852-1740. Identify all comments with the docket number listed in the notice of availability that publishes in the *Federal Register*.

For questions about this document regarding CDRH-regulated devices, contact the Office of Product Evaluation and Quality, Office of Regulatory Programs, Division of Regulatory Programs 1 (Submission Support) at 301-796-5640. For questions about this document regarding CBER-regulated devices, contact the Office of Communication, Outreach, and Development (OCOD) at 1-800-835-4709 or 240-402-8010, or by email at [ocod@fda.hhs.gov](mailto:ocod@fda.hhs.gov).

**U.S. Department of Health and Human Services**

**Food and Drug Administration Center for Devices and Radiological Health Center for Biologics Evaluation and Research**

**Preface**

# Additional Copies CDRH

Additional copies are available from the Internet. You may also send an email request to [CDRH-](mailto:CDRH-Guidance@fda.hhs.gov) [Guidance@fda.hhs.gov](mailto:CDRH-Guidance@fda.hhs.gov) to receive a copy of the guidance. Please include the document number GUI00020006 and complete title of the guidance in the request.

# CBER

Additional copies are available from the Center for Biologics Evaluation and Research (CBER), Office of Communication, Outreach, and Development (OCOD), 10903 New Hampshire Ave., Bldg. 71, Room 3128, Silver Spring, MD 20993-0002, or by calling 1-800-835-4709 or 240-402- 8010, by email, [ocod@fda.hhs.gov](mailto:ocod@fda.hhs.gov), or from the Internet at [https://www.fda.gov/vaccines-blood-](https://www.fda.gov/vaccines-blood-biologics/guidance-compliance-regulatory-information-biologics/biologics-guidances) [biologics/guidance-compliance-regulatory-information-biologics/biologics-guidances](https://www.fda.gov/vaccines-blood-biologics/guidance-compliance-regulatory-information-biologics/biologics-guidances).

**Table of Contents**

1. [Introduction](#_bookmark0) [1](#_bookmark0)
2. [Background](#_bookmark2) [2](#_bookmark2)
   1. [The 510(k) Process](#_bookmark3) [2](#_bookmark3)
   2. [510(k) Modernization](#_bookmark7) [3](#_bookmark7)
3. [Scope](#_bookmark11) [4](#_bookmark11)
4. [How to use this guidance](#_bookmark12) [4](#_bookmark12)
5. [Best practices for selecting a predicate device](#_bookmark21) [6](#_bookmark21)
   1. [Predicate devices cleared using well-established methods](#_bookmark23) [7](#_bookmark23)
   2. [Predicate devices meet or exceed expected safety and performance](#_bookmark30) [8](#_bookmark30)
   3. [Predicate devices without unmitigated use-related or design-related safety issues](#_bookmark35) [9](#_bookmark35)
   4. [Predicate devices without an associated design-related recall](#_bookmark41) [10](#_bookmark41)
6. [Improving the Transparency of Predicate Devices](#_bookmark46) [11](#_bookmark46)
7. [Examples](#_bookmark49) [12](#_bookmark49)

1 **Best Practices for Selecting a Predicate**

2 **Device to Support a Premarket**

3 **Notification [510(k)] Submission**

4

5 **Draft Guidance for Industry and**

6 **Food and Drug Administration Staff**

7

***This draft guidance, when finalized, will represent the current thinking of the Food and Drug Administration (FDA or Agency) on this topic. It does not establish any rights for any person and is not binding on FDA or the public. You can use an alternative approach if it satisfies the requirements of the applicable statutes and regulations. To discuss an alternative approach, contact the FDA staff or Office responsible for this guidance as listed on the title page.***

8

9

10

11

12

13

14

# I. Introduction

1. FDA developed this document to provide guidance to industry and FDA staff about best
2. practices in selecting a predicate device for premarket notification [510(k)] submissions.
3. Specifically, this guidance recommends four (4) best practices to employ when selecting a
4. predicate device used to support a 510(k) submission. The recommendations provided in this
5. guidance are not intended to propose any changes to applicable statutory and regulatory
6. standards, such as how FDA evaluates substantial equivalence, or the applicable requirements,
7. including the requirement for valid scientific evidence. FDA developed this guidance to improve
8. the predictability, consistency, and transparency of the 510(k) premarket review process. This
9. [guidance and associated recommendations are consistent with and are intended to be used in](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k)
10. conjunction with the FDA guidance “The 510(k) Program: Evaluating Substantial Equivalence in
11. Premarket Notifications [510(k)]”1 (hereinafter, 510(k) Program Guidance) and other relevant
12. FDA guidances on 510(k) submissions. 28
13. In general, FDA’s guidance documents do not establish legally enforceable responsibilities.
14. Instead, guidances describe the Agency’s current thinking on a topic and should be viewed only
15. as recommendations, unless specific regulatory or statutory requirements are cited. The use of
16. the word *should* in Agency guidances means that something is suggested or recommended, but
17. not required.

34

1 Available at [https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k) [evaluating-substantial-equivalence-premarket-notifications-510k](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k).

# II. Background

## A. The 510(k) Process

1. The framework under which FDA regulates medical devices was put into place when Congress
2. enacted the Medical Device Amendments (Pub. L. 94-295) to the Federal Food, Drug, and
3. Cosmetic Act (FD&C Act) on May 28, 1976. Under section 510(k) of the FD&C Act, a
4. manufacturer must submit a premarket notification (often referred to as a 510(k)) to FDA at least
5. 90 days before introducing, or delivering for introduction, a device into interstate commerce for
6. commercial distribution so the Agency can determine whether or not the device meets the criteria
7. for market clearance (sections 510(k), 510(n), and 513(i) of the FD&C Act). A 510(k) is required
8. for devices intended for human use, for which a premarket approval application (PMA) is not
9. required, unless the device is exempt from the 510(k) requirements of the FD&C Act and does
10. not exceed the relevant limitations of exemptions in the device classification regulations. 47
11. A 510(k) is a marketing submission made by a manufacturer to FDA to demonstrate that the
12. device to be marketed is substantially equivalent to a “predicate device” (section 513(i) of the
13. FD&C Act and 21 CFR 807.92(a)(3)-(6)). Substantial equivalence is rooted in a comparison
14. between the “new device”2 and predicate device(s).3 52
15. The Agency bases its decision on whether the device is substantially equivalent (SE) to a
16. predicate device using the statutory criteria in section 513(i) of the FD&C Act. For FDA to find a
17. new device SE to a predicate device, FDA must first find that the new device and predicate
18. device have the same intended use. FDA must then find that the new device and predicate device
19. have the same technological characteristics, or if they do not, that the different technological
20. characteristics4 of the new device do not raise different questions of safety and effectiveness and
21. that the new device is as safe and effective as a predicate device. FDA conducts this evaluation
22. by reviewing the proposed scientific methods for evaluating new/different technological
23. characteristics’ effects on safety and effectiveness and accompanying performance data to
24. determine whether the methods are acceptable and whether the data demonstrates SE. A new
25. device requiring premarket notification cannot be introduced into interstate commerce for

2 For purposes of this guidance, a “new device” means a device within the meaning of section 201(h) of the FD&C Act that is not legally marketed. It can be either a completely new device (i.e., one that has not received FDA’s marketing authorization) or a modification of a legally marketed device that would require a new 510(k).

3 A predicate device is a legally marketed device. Under 21 CFR 807.92(a)(3), a legally marketed device to which a new device may be compared for a determination regarding substantial equivalence is a device that was legally marketed prior to May 28, 1976, or a device which has been reclassified from class III to class II or I, or a device which has been found to be substantially equivalent through the 510(k) premarket notification process. Moreover, “[a] device may not be found to be substantially equivalent to a predicate device that has been removed from the market at the initiative of [FDA] or that has been determined to be misbranded or adulterated by a judicial officer.” Section 513(i)(2) of the FD&C Act.

4 For purposes of an SE determination, “‘different technological characteristics’ means, with respect to a device being compared to a predicate device, that there is a significant change in the materials, design, energy source, or other features of the device from those of the predicate device.” Section 513(i)(1)(B) of the FD&C Act.

1. commercial distribution until FDA issues an order stating that the device has been determined to
2. be SE (s[ection 513(f)(1) of the FD&C Act).5](https://www.fda.gov/about-fda/cdrh-reports/medical-device-safety-action-plan-protecting-patients-promoting-public-health) [66](https://www.fda.gov/about-fda/cdrh-reports/medical-device-safety-action-plan-protecting-patients-promoting-public-health)

## [B. 510(k) Modernization](https://www.fda.gov/about-fda/cdrh-reports/medical-device-safety-action-plan-protecting-patients-promoting-public-health)

1. In April 2018, CDRH issued the Medical Device Safety Action Plan: Protecting Patients,
2. Promoting Public Health6 (herein referred to as the “Safety Action Plan”) to communicate
3. CDRH’s vision for modernizing measures to improve the safety of medical devices while
4. continuing to create more efficient pathways to bring critical devices to patients. The Safety
5. Action Plan describes the efforts underway to modernize the 510(k) program. 73
6. In November 2018, FDA announced transformative new steps to modernize FDA’s 510(k)
7. program to advance the review of the safety and effectiveness of medical devices. In connection
8. with this announcement, FDA also requested public feedback on these steps to continue to
9. modernize the framework for 510(k) review while promoting innovation and improving safety
10. by driving innovators toward reliance on more modern predicate devices or objective
11. performance criteria when they seek to bring new devices to the market and ultimately to
12. patients. FDA indicated that it is looking at ways to promote the use of more recent predicates
13. because it believes that newer devices should be compared to the benefits and risks of more
14. modern technology.

83

1. To advance these goals, FDA discussed several potential options for 510(k) modernization. The
2. statement discussed that the Agency considered making public on its website those cleared
3. devices that demonstrated substantial equivalence to older predicate devices. FDA also
4. considered focusing on predicates that were more than ten (10) years old as a starting point, so
5. the public was made aware of those technologies. FDA’s goal in focusing on older predicates
6. was to encourage manufacturers to continually offer patients devices with the latest
7. improvements and advances. FDA issued a public notice on January 22, 20197 on FDA’s website
8. that requested public comment on this proposal. 92
9. FDA reviewed all comments submitted to the docket and acknowledges that the initial proposal
10. of focusing only on older predicates may not optimally promote safer and more effective
11. devices. For example, if selecting a predicate for an implant, older devices may potentially have
12. long-term safety and effectiveness data that establishes a history of safe and effective use.
13. Conversely, when selecting a predicate for a device that includes software, a more recently

5 Under section 510(k) of the FD&C Act, premarket notification is required for devices that are not subject to a premarket approval application, unless the device is exempt from the 510(k) requirements of the FD&C Act and does not exceed the limitations of exemptions for each of the device classification regulations (e.g., 21 CFR Parts 862-892).

6 Available at [https://www.fda.gov/about-fda/cdrh-reports/medical-device-safety-action-plan-protecting-patients-](https://www.fda.gov/about-fda/cdrh-reports/medical-device-safety-action-plan-protecting-patients-promoting-public-health) [promoting-public-health](https://www.fda.gov/about-fda/cdrh-reports/medical-device-safety-action-plan-protecting-patients-promoting-public-health).

7 Available at https://wayback.archive- it.org/7993/20190206202131/https://[www.fda.gov/AboutFDA/CentersOffices/OfficeofMedicalProductsandTobacco](http://www.fda.gov/AboutFDA/CentersOffices/OfficeofMedicalProductsandTobacco)

/CDRH/CDRHReports/ucm604500.htm. Public comments submitted can be searched under the docket FDA-2018- N-4751, available at <https://www.regulations.gov/docket/FDA-2018-N-4751/comments>.

1. cleared device could include modern safety features due to rapid technological advances that
2. affect cybersecurity, interoperability, and modern software architectures. 100
3. After considering the docket comments, FDA believes that it may be more appropriate to
4. modernize the 510(k) process with respect to the use of predicate devices by focusing on
5. utilizing best practices when selecting a predicate device rather than just their age. Therefore,
6. FDA is issuing this draft guidance to propose ways to encourage the use of best practices when
7. selecting a predicate device. 106
8. FDA developed this draft guidance to propose factors for consideration as best practices for
9. choosing a predicate device. These best practices include consideration of the characteristics of
10. predicate devices rather than focusing on the age of the predicate. FDA believes that this will
11. encourage the evolution of safer and more effective medical devices in the 510(k) program over
12. time. Additionally, FDA believes that identification of the characteristics of predicate devices
13. used to support a 510(k) submission in the accompanying 510(k) Summary may provide
14. additional transparency to the public for devices subject to 510(k) requirements.8 114

# III. Scope

1. This guidance provides recommendations to industry and FDA staff about the best practices of
2. choosing a predicate device for a 510(k) submission. This guidance is intended to be used in
3. conjunction with the 510(k) Program Guidance.9 The recommendations provided in this
4. guidance are not intended to propose any changes to applicable statutory and regulatory
5. standards, such as how FDA evaluates substantial equivalence, or the applicable requirements,
6. including the requirement for valid scientific evidence*.* FDA developed this guidance to improve
7. the predictability, consistency, and transparency of the 510(k) premarket review process. 123
8. This guidance is also not intended to supplant existing device-specific guidance but may cover
9. broader areas not addressed in device-specific guidances. 126

# IV. How to use this guidance

1. This guidance is intended to guide submitters through the best practices in selecting a predicate
2. device for a 510(k) submission. This guidance is intended to be used while a submitter is
3. preparing their 510(k) submission to assist with the identification of potential predicate device(s)
4. to support their device’s substantial equivalence to a legally marketed device. Based on FDA’s
5. experience in reviewing 510(k) submissions, the Agency is aware that many submitters include

8 Consistent with the [510(k) Program Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k) and as specified in 21 CFR 807.92(a)(6), the 510(k) Summary shall contain the following information: “If the device has the same technological characteristics (i.e., design, material, chemical composition, energy source) as the predicate device identified in [21 CFR 807.92(a)(3)], a summary of the technological characteristics of the subject device in comparison to those of the predicate device. If the device has different technological characteristics from the predicate device, a summary of how the technological characteristics of the device compare to a legally marketed device identified in [21 CFR 807.92(a)(3)].”

9 Available at [https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k)

1. in their 510(k) submission a completed 510(k) flowchart with a discussion describing why the
2. submitter believes their device is substantially equivalent to the predicate device.10 135
3. When considering the selection of predicate devices during 510(k) submission preparation,
4. submitters should consider the list of legally marketed devices that they believe have the same
5. intended use as the subject device and when any differences in technological characteristics do
6. not raise different questions about safety and effectiveness, hereafter referred to as a “valid
7. predicate device.”11, 12 FDA recommends narrowing this list of valid predicate device(s) to the
8. predicate device13 identified by the submitter to support the 510(k) submission using the best
9. practices outlined in Section V of this guidance, in conjunction with the 510(k) Program
10. Guidance.14 A visual representation of this concept is illustrated in Figure 1 below. 144

145

10 For example, FDA has guidance on the “[Format for Traditional and Abbreviated 510(k)s](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/format-traditional-and-abbreviated-510ks),” available at [https://www.fda.gov/regulatory-information/search-fda-guidance-documents/format-traditional-and-abbreviated-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/format-traditional-and-abbreviated-510ks) [510ks](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/format-traditional-and-abbreviated-510ks). Submitters could include such an assessment in Section 12 (Substantial Equivalence Discussion) of their 510(k) submission.

11 Consistent with sections 510(k), 510(n), and 513(i) of the FD&C Act and the [510(k) Program Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k), while submitters propose a predicate device in their 510(k) submission, FDA determines whether a subject and predicate device are substantially equivalent. This determination includes whether a valid predicate device exists for the subject device.

12 Consistent with the [510(k) Program Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k), if FDA has established special controls applicable to the device type, the 510(k) would also need to demonstrate that the proposed device meets the relevant special controls for the device to be classified into class II.

13 Consistent with the [510(k) Program Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k), a submitter may use multiple predicate devices to help demonstrate substantial equivalence in certain circumstances. Submitters sometimes choose to do this when combining features from two or more predicate devices with the same intended use into a single new device, when seeking to market a device with more than one intended use, or when seeking more than one indication for use under the same intended use. Additionally, while FDA does not consider reference devices to be predicate devices, reference devices can be used to support a 510(k) submission beyond Decision 4 in the 510(k) flowchart (See Appendix A in the [510(k)](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k) [Program Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k)). For example, reference devices can be used to support scientific methodology or standard reference values at Decision 5a in the 510(k) flowchart.

14 Available at [https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k)

146

147

148

149

150

151

152

153

154

155

156

[157](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm)

158

159

160

161


162

163

164

165

166

**Figure 1.** Visual depiction of the relevant terminology used in this guidance.

FDA recommends the submitter include within their 510(k) submission how they used the best practices identified in this guidance in selecting the predicate device(s) used to support the 510(k) submission. For example, if a valid predicate device consistent with the best practices identified in this guidance is not available, FDA recommends describing in the 510(k) submission how any known concerns with the valid predicate device have been mitigated with the subject device (e.g., design features, performance testing). FDA also recommends that the submitter summarize how the best practices were utilized in the selection of the predicate device used to support the 510(k) submission in the 510(k) Summary (See Section VI of this guidance). These recommendations are intended to aid the submitter in selecting a predicate for [their device](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm) [and help provide](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm) additional transparency to the public in the 510(k) summary if the 510(k) submission is cleared by FDA.

# Best practices for selecting a predicate device

FDA identifies all devices cleared through the 510(k) process in the publicly available FDA 510(k) Premarket Notification Database.15 This online database is updated monthly by FDA. Most submitters likely start with basic administrative information to identify valid predicate device(s), including but not limited to the:

* + Trade names of similar devices;

|  |  |
| --- | --- |
| 167 | * Manufacturer(s) of similar devices; |
| 168 | * 510(k) numbers for similar devices; and |
| 169 | * Searching of classification information (e.g., product codes, classification regulation) for |
| 170 | similar devices. |
| 171 |  |

15 Available at <https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm>.

1. Once legally marketed devices have been identified, FDA recommends reviewing the publicly-
2. available 510(k) Summary16 and Indications for Use documents for each device being considered
3. by the submitter as a valid predicate device. In addition to these basic administrative items, FDA
4. recommends submitters apply the best practices identified below when selecting a predicate
5. device to support the 510(k) submission. 177

## A. Predicate devices cleared using well­established methods

1. FDA recommends selecting a valid predicate device that was cleared using well-established
2. methods.17 These methods include those from a currently FDA-recognized voluntary consensus
3. standard,18 an FDA guidance document,19 a qualified medical device development tool
4. (MDDT),20 or a widely available and accepted method published in the public domain or
5. scientific literature for the context of use, or found acceptable through the submitter’s own
6. previous premarket submission. FDA recommends prioritizing predicate devices with methods
7. developed within a consensus environment, and those subject to public comment or peer review.
8. FDA believes that when selecting a valid predicate device, submitters should consider how much
9. information is available regarding the test method(s) used in support of the predicate device’s
10. 510(k) clearance and whether those methods continue to be appropriate for evaluating the subject
11. device. For example, voluntary consensus standards periodically undergo revisions and the
12. methods used to evaluate devices can change with both industry and FDA experience with a
13. device.

192

1. Once the submitter has identified a list of valid predicate devices, FDA recommends conducting
2. a search of the nonclinical tests submitted, referenced, or relied on in the 510(k) submission21 to
3. support a determination of substantial equivalence. For example, when selecting between two
4. similar valid predicate devices, where one identified performing testing using FDA guidance and

16 As specified in 21 CFR 807.87(h), a 510(k) Statement as described in 21 CFR 807.93 may be provided by the submitter in lieu of a 510(k) Summary. However, in order to facilitate transparency, FDA encourages all submitters to utilize the 510(k) Summary option.

17 FDA acknowledges this information may not always be publicly available for a predicate device, especially for those that were not recently cleared.

18 For the current edition of FDA-recognized standard(s), see the [Recognized Consensus Standards Database](http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfStandards/search.cfm), available at <http://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfStandards/search.cfm>. For more information regarding use of consensus standards in regulatory submissions, refer to the FDA guidance titled “[Appropriate Use](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/appropriate-use-voluntary-consensus-standards-premarket-submissions-medical-devices) [of Voluntary Consensus Standards in Premarket Submissions for Medical Devices](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/appropriate-use-voluntary-consensus-standards-premarket-submissions-medical-devices),” available at [https://www.fda.gov/regulatory-information/search-fda-guidance-documents/appropriate-use-voluntary-consensus-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/appropriate-use-voluntary-consensus-standards-premarket-submissions-medical-devices) [standards-premarket-submissions-medical-devices](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/appropriate-use-voluntary-consensus-standards-premarket-submissions-medical-devices).

19 A list of FDA guidance documents is available on FDA’s website at [https://www.fda.gov/regulatory-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents) [information/search-fda-guidance-documents](https://www.fda.gov/regulatory-information/search-fda-guidance-documents).

20 A list of qualified MDDTs is available on FDA’s website at [https://www.fda.gov/medical-devices/medical-](https://www.fda.gov/medical-devices/medical-device-development-tools-mddt) [device-development-tools-mddt](https://www.fda.gov/medical-devices/medical-device-development-tools-mddt).

21 In accordance with 21 CFR 807.92(a)(3), the 510(k) Summary must identify the predicate relied upon for a substantial equivalence determination. In accordance with 21 CFR 807.92, FDA describes the requirements and recommendations of the content to be included in a 510(k) Summary in Appendices B and C of the [510(k) Program](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k) [Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k). 510(k) Summaries for devices that have been cleared for marketing through the FDA can be found in the [510(k) Premarket Notification Database](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm) on the FDA website at <https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm>.

1. FDA recognized voluntary consensus standards while the other did not, FDA recommends the
2. submitter select the valid predicate device in which testing relied on these well-established
3. methods. FDA considers it a best practice to select a predicate that was cleared using well-
4. established methods, as this will continue to advance the 510(k) Program, by encouraging the
5. evolution of safer and more effective medical devices in the 510(k) program over time, and
6. ensure that the subject device is evaluated using updated scientific methods whenever possible. 203

## B. Predicate devices meet or exceed expected safety and

1. **performance**
2. FDA considers it a best practice to select a valid predicate device that continues to perform
3. safely and as intended by the manufacturer during use in its intended environment of use
4. whenever possible. FDA recommends selecting a valid predicate device after considering how
5. any reported medical device-related adverse events, malfunctions, or deaths may have a role in
6. the safety and effectiveness of the device. New information about a device’s safety and/or
7. effectiveness, including unanticipated adverse events, may become available once the device is
8. more widely distributed and used commercially. Also, subsequent changes made to the device,
9. including material changes, or its manufacturing process may lead to unanticipated effects that
10. cannot be comprehensively captured during premarket review.22 This new information may
11. include, but is not limited to, a newly recognized type of adverse event associated with a medical
12. device, an increase in the severity or frequency of a known adverse event, new product-product
13. interactions, or device malfunctions. 218
14. Once the submitter has identified a list of valid predicate devices, [FDA recommends conducting](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/search.cfm)
15. [a search for any reported injury, death](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/search.cfm)s[,](#_bookmark32) [or malfunctions using](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmdr/search.cfm) the following FDA databases:
16.  Manufacturer and User Facility Device Experience (MAUDE) Database;23
17.  Medical Device Reporting (MDR) Database;24 and
18.  MedSun Reports Database.25 224
19. FDA recommends searching each of the above databases for any reports of unexpected injury,
20. deaths, or malfunctions associated with the available valid predicate devices. For example, when
21. selecting a predicate device for an infusion pump, if the database search reveals a high frequency
22. of reports of battery failures related to the predicate device that resulted in serious injuries to the
23. operator, such events could suggest fundamental design issues with this valid predicate device
24. and FDA recommends selection of a different valid predicate device for the 510(k) submission

22 The regulatory criteria for when a premarket notification submission is required for a change to an existing device are outlined in 21 CFR 807.81(a)(3). For more information regarding when a premarket modification submission is required refer to the FDA guidance titled, “[Deciding When to Submit a 510(k) for a Change to an Existing Device](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-510k-change-existing-device),” available at [https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-510k-change-existing-device) [510k-change-existing-device](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-510k-change-existing-device), and the FDA guidance titled, “[Deciding When to Submit a 510(k) for a Software](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-510k-software-change-existing-device) [Change to an Existing Device](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-510k-software-change-existing-device),” available at [https://www.fda.gov/regulatory-information/search-fda-guidance-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-510k-software-change-existing-device) [documents/deciding-when-submit-510k-software-change-existing-device](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/deciding-when-submit-510k-software-change-existing-device).

23 Available at <https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/search.cfm>.

24 Available at <https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmdr/search.cfm>.

25 Available at <https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/medsun/searchreporttext.cfm>.

|  |  |
| --- | --- |
| 231 | whenever possible. If another valid predicate device is not available, FDA recommends that the |
| 232 | submitter describe in the 510(k) submission how the subject device mitigates the known |
| 233 | concerns with the predicate device used to support the 510(k) submission. |
| 234 |  |
| 235 | **C. Predicate devices without unmitigated use­related or design­** |
| 236 | **related safety issues** |
| 237 | FDA recommends selecting a valid predicate device that does not have unmitigated use-related |
| 238 | or design-related safety issues, including consideration of emerging signals or safety |
| 239 | communications.26 New information about a device’s safety and/or effectiveness can become |
| 240 | available once the device is more widely distributed and used. This new information could |
| [241](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/public-notification-emerging-postmarket-medical-device-signals-emerging-signals) | represent a signal and may include information related to [device malfunctions or patient injuries](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/public-notification-emerging-postmarket-medical-device-signals-emerging-signals) |
| 242 | potentially related to improper device use or design. |
| 244 | Consistent with the FDA guidance “Public Notification of Emerging Postmarket Medical Device |
| 245 | Signals (“Emerging Signals”),”27 an emerging signal is new information about a device that |
| 246 | supports a new causal association or a new aspect of a known association between a device and |
| 247 | an adverse event or set of adverse events and for which FDA has conducted an initial evaluation |
| 248 | and determined that the information has the potential to impact patient management decisions |
| 249 | and/or the known benefit-risk profile of the device. An emerging signal may be associated with |
| [250](https://www.fda.gov/vaccines-blood-biologics/safety-availability-biologics) | one product from one manufacturer, one type of product [or similar products from multiple](https://www.fda.gov/vaccines-blood-biologics/safety-availability-biologics) |
| 251 | manufacturers, or multiple different product types from multiple different manufacturers (e.g., |
| 252 | materials issues). Information about emerging signals and safety communications is available on |
| 253 | the Medical Device Safety28 and CBER Safety & Availability (Biologics)29 websites. FDA |
| 254 | [recommends reviewing](https://www.fda.gov/medical-devices/medical-device-safety) any [safety signals, emerging signals,](https://www.fda.gov/vaccines-blood-biologics/safety-availability-biologics) or other safety information |
| 255 | available prior to selecting a valid predicate device to support the 510(k) submission. |
| 257 | Once the submitter has identified a list of valid predicate devices, FDA recommends conducting |
| 258 | a search of the Medical Device Safety and CBER Safety & Availability (Biologics) websites to |
| 259 | assess whether any of the valid predicate devices have an associated use-related or design-related |
| 260 | safety issue. For example, a signal was reported for duodenoscopes describing challenges in the |
| 261 | adequacy of reprocessing instructions that resulted in a potential for disease transmission.30 This |
| 262 | signal resulted in an effort to replace traditionally reprocessed duodenoscopes with |
| 263 | duodenoscopes which had innovative designs to enhance safety, including designs with |
| 264 | disposable caps or distal ends. FDA considers it a best practice to select a valid predicate device |
| 265 | that is not associated with emerging signals or safety communications that relate to unmitigated |
| 266  267 | use-related or design-related safety issues whenever possible. |
|  | 26 Available at <https://www.fda.gov/medical-devices/medical-device-safety>.  27 Available at [https://www.fda.gov/regulatory-information/search-fda-guidance-documents/public-notification-](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/public-notification-emerging-postmarket-medical-device-signals-emerging-signals) |
|  | [emerging-postmarket-medical-device-signals-emerging-signals](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/public-notification-emerging-postmarket-medical-device-signals-emerging-signals).  28 Available at <https://www.fda.gov/medical-devices/medical-device-safety>.  29 Available at <https://www.fda.gov/vaccines-blood-biologics/safety-availability-biologics>.  30 Available at [https://www.fda.gov/medical-devices/safety-communications/use-duodenoscopes-innovative-](https://www.fda.gov/medical-devices/safety-communications/use-duodenoscopes-innovative-designs-enhance-safety-fda-safety-communication) |
|  | [designs-enhance-safety-fda-safety-communication](https://www.fda.gov/medical-devices/safety-communications/use-duodenoscopes-innovative-designs-enhance-safety-fda-safety-communication). |

243

256

## D. Predicate devices without an associated design­related recall

1. FDA recommends selecting a valid predicate device that has not been subject to a design-related
2. recall.31 Recalls are typically voluntary actions taken by a manufacturer or may be requested by
3. FDA to correct or remove a violative product from the market.32 A violative product is one in
4. violation of the laws that FDA administers and against which FDA would initiate legal action.
5. Recalls can occur due to design defects, manufacturing defects, or labeling defects. 274
6. Design-related recalls can indicate a fundamental flaw with the design of the device as cleared
7. and commercially distributed. Design controls under 21 CFR 820.30 include a framework that
8. requires manufacturers subject to these requirements to establish and maintain procedures to
9. control the design of the device in order to ensure that specified design requirements are met.33,34
10. When a design-related recall has been conducted for a device, adequate design control
11. procedures, including but not limited to design input, output, verification, validation, and transfer
12. may not have been adequately implemented through the design process. In some instances, the
13. underlying root cause of the design related issues identified as part of a design-related recall may
14. not be available or a correction of these design-related issues may not be possible. Further,
15. although the methods and performance data provided in the 510(k) submission for the valid
16. predicate device subject to a subsequent design-related recall were sufficient to support a
17. substantial equivalence determination at that time of 510(k) clearance, utilization of such a valid
18. [predicate device may not be](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm) ideal to use for future 510(k) submissions. 288
19. Once the submitter has identified a list of valid predicate devices, FDA recommends conducting
20. a search of the Medical Device Recalls Database to assess whether any of the valid predicate
21. devices have an associated recall. For example, the recall of a coronary catheter tip for fracture
22. could be associated with a change in the manufacturing process or with the materials used in the
23. design of the catheter. If a recall is associated with the materials used in the catheter, this could
24. be considered a design-related recall when it is due to an inadequacy of that material to meet user

31 The [Medical Device Recalls Database](https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm) is available at <https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm>, and includes recalls classified since November 2002.

32 While FDA focuses on voluntary recalls conducted under 21 CFR Part 7, FDA may, after providing the appropriate person with an opportunity to consult with the Agency, also require that a manufacturer recall their device when the criteria are met under 21 CFR Part 810.

33 For devices subject to 510(k) requirements, design controls apply to class II devices and those class I devices listed in 21 CFR 820.30(a)(2).

34 On February 23, 2022, FDA proposed to amend the device Quality System Regulation, 21 CFR Part 820, to align more closely with international consensus standards for devices ([87 FR 10119](https://www.federalregister.gov/documents/2022/02/23/2022-03227/medical-devices-quality-system-regulation-amendments); available at [https://www.federalregister.gov/documents/2022/02/23/2022-03227/medical-devices-quality-system-regulation-](https://www.federalregister.gov/documents/2022/02/23/2022-03227/medical-devices-quality-system-regulation-amendments) [amendments](https://www.federalregister.gov/documents/2022/02/23/2022-03227/medical-devices-quality-system-regulation-amendments)). Specifically, FDA proposed to withdraw the majority of the current requirements in Part 820 and instead incorporate by reference the 2016 edition of the International Organization for Standardization (ISO) 13485, Medical devices – Quality management systems for regulatory purposes, in Part 820. As stated in that proposed rule, the requirements in ISO 13485 are, when taken in totality, substantially similar to the requirements of the current Part 820, providing a similar level of assurance in a firm’s quality management system and ability to consistently manufacture devices that are safe and effective and otherwise in compliance with the FD&C Act. FDA intends to finalize this proposed rule expeditiously. When the final rule takes effect, FDA will also update the references to provisions in 21 CFR Part 820 in this guidance to be consistent with that rule.

1. needs and intended uses. The material used in the design of this catheter and the testing
2. conducted on the predicate device did not adequately mitigate against the risk of tip fracture,
3. resulting in a design-related recall. FDA considers it a best practice to select a valid predicate
4. device that is not associated with a design-related recall whenever possible. 299

# VI. Improving the Transparency of Predicate Devices

1. The 510(k) Summary is a document that provides an adequate summary of any information
2. respecting safety and effectiveness and must include all the elements identified in 21 CFR
3. 807.92.35 A 510(k) Summary must be in sufficient detail to provide an understanding of the basis
4. for a determination of substantial equivalence (21 CFR 807.92(a)). In Appendix B of the 510(k)
5. Program Guidance, FDA describes the requirements of the content to be included in a 510(k)
6. Summary, in accordance with 21 CFR 807.92, and provides recommendations on the
7. information to be included in a 510(k) Summary to ensure compliance with 21 CFR 807.92 and
8. consistency in the level of information conveyed and captured in the 510(k) Summaries that are
9. available to the public on FDA’s website. 310
10. In an effort to improve the transparency and predictability of the 510(k) program and to ensure
11. that the 510(k) Summary reflects the information provided in a 510(k) submission to support a
12. substantial equivalence determination, FDA stated in the 510(k) Program Guidance that the
13. Agency intends to verify the accuracy and completeness of the information included in a 510(k)
14. Summary.

316

1. Although the 510(k) Summary is a document drafted by the submitter and is included in the
2. 510(k), revisions to the 510(k) Summary may be necessary to accurately reflect FDA’s decision-
3. making process. As stated in the 510(k) Program Guidance, and consistent with 21 CFR
4. 807.92(b)(1), 510(k) Summaries shall include a brief discussion of the nonclinical tests
5. submitted, referenced, or relied on in the premarket notification submission for a determination
6. of substantial equivalence. 323
7. FDA recommends that submitters include a narrative explaining their selection of the predicate
8. device(s) used in support of the 510(k) submission in their draft 510(k) Summary submitted with
9. their original 510(k).36 FDA recommends this narrative include a discussion of how the best
10. practices described in Section V of this guidance were used to select the predicate device(s)
11. proposed for use in the 510(k) submission. This recommendation is intended to promote
12. transparency to the public regarding the process of selecting a predicate device using these best
13. practices.

331

1. When a submitter cannot identify a valid predicate device(s) that is consistent with any of the
2. best practices discussed in Section V of this guidance, FDA recommends that the submitter

35 As specified in 21 CFR 807.87(h), a 510(k) Statement as described in 21 CFR 807.93 may be provided in lieu of a 510(k) Summary. However, in order to facilitate transparency, FDA encourages all submitters to utilize the 510(k) Summary option.

36 As described in Appendix C of the [510(k) Program Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k), this information is provided in the Comparison of Technological Characteristics with the Predicate Device of the 510(k) Summary.

1. include a statement in their 510(k) Summary that a valid predicate that is consistent with the best
2. practices was not available. FDA recommends that the submitter also use the Performance Data
3. Section of the 510(k) Summary to describe the ways performance testing was conducted to
4. address any known safety or effectiveness concerns with the predicate device used to support the
5. 510(k) submission.37

339

# VII. Examples

1. The following are illustrative examples that are intended to exemplify how the best practices
2. identified in Section V of this guidance for selecting a valid predicate device can be used. These
3. examples do not necessarily account for every possible detail, risk, or consideration that a
4. submitter should consider when selecting the predicate device used in support of the 510(k)
5. submission. These examples also include different formats that could be used depending on the
6. number of valid predicate devices available for use to support the 510(k) submission. 347

348 

1. A submitter is preparing a 510(k) submission for a coronary guidewire, Guidewire X. The
2. submitter identified four valid predicate devices, all of which have the same intended use as
3. Guidewire X and any differences in technological characteristics do not raise different questions
4. of safety and effectiveness. The submitter included the following table in their 510(k)
5. submission, along with their rationale for selecting Predicate 4 as the predicate device used to
6. support their 510(k) submission: 355


37 Section 518A of the FD&C Act directs FDA to establish a program to routinely and systematically assess information regarding device recalls, and to use that information to proactively identify strategies for mitigating health risks presented by defective or unsafe devices. Consistent with the [510(k) Program Guidance](https://www.fda.gov/regulatory-information/search-fda-guidance-documents/510k-program-evaluating-substantial-equivalence-premarket-notifications-510k), FDA believes that providing greater transparency on recalled devices is one way to help achieve this directive.

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Valid Predicate Device** | **A – Well- established methods** | **B – Meets or exceeds expected predicate performance** | **C –**  **Unmitigated use-related or design-related safety issues** | **D – Associated design-related recall** |
| 1 | Used internal methods that are not widely available and accepted | High frequency of fractures reported in MDRs/MedSun reports | Safety communication found on FDA’s website | No design- related recall identified |
| 2 | Used relevant methods that were published in the public domain | Expected frequency of reported adverse events | No known unmitigated use- related or design- related safety issues | Design-related recall identified in FDA’s database |
| 3 | Used outdated methods in a subsequently superseded FDA  guidance document | Expected frequency of reported adverse events | No known unmitigated use- related or design- related safety  issues | No design- related recall identified |
| 4 | Used updated methods from current FDA guidance document | Expected frequency of reported adverse events | No known unmitigated use- related or design- related safety issues | No design- related recall identified |


356

1. In their draft 510(k) Summary, the submitter includes a brief narrative describing the above
2. selection process in the proposed 510(k) Summary. The submitter’s draft 510(k) Summary also
3. includes a discussion that the selected predicate used well-established methods from a current
4. FDA guidance document, discusses the frequency of reported adverse events, and states that
5. there are no known unmitigated use-related or design-related safety issues or design-related
6. recalls.

363

364 

1. A submitter is preparing a 510(k) submission for a bone sonometer, Bone Sonometer X. The
2. submitter identified only one valid predicate device, which has the same intended use as bone
3. sonometer X and any differences in technological characteristics do not raise different questions
4. of safety and effectiveness. The valid predicate device, Predicate 1, used the currently FDA-
5. recognized versions of applicable consensus standards, has an expected frequency of reported
6. events, had no known unmitigated use-related or design-related safety issues before submission
7. of the device-related recall, but has been associated with a design-related recall. 372
8. The submitter referenced Predicate 1 as their predicate device in their 510(k) submission, along
9. with a statement that Predicate 1 was the only valid predicate device that could be identified. The
10. submitter also described the ways performance testing was conducted to address the safety
11. concerns relevant to the design-related recall associated with Predicate 1 and the measures taken
12. to mitigate those safety concerns in the subject device. 378
13. In their draft 510(k) Summary, the submitter identified that the predicate device used to support
14. the 510(k) submission has been the subject of a design-related recall, but also included a brief
15. narrative describing the selection process in the proposed 510(k) Summary. The sponsor also
16. included a summary of how their performance testing provided in the 510(k) addressed the safety
17. concerns relevant to Predicate 1’s design-related recall in the Performance Data section of the
18. draft 510(k) Summary. 385

386 

1. A submitter is preparing a 510(k) submission for an intervertebral fusion device (IFD), IFD X.
2. The submitter identified two valid predicate devices, both of which have the same intended use
3. as IFD X and any differences in technological characteristics do not raise different questions of
4. safety and effectiveness. Predicate 1 has been on the market for 15 years and Predicate 2 has
5. been on the market for 3 years and both devices are still in clinical use. The submitter included
6. the following table in their 510(k) submission, along with their rationale describing that while
7. Predicate 2 also uses the best practices for selecting a predicate device, Predicate 1 was selected
8. as the predicate device used to support the 510(k) submission because it has a well-established
9. safety profile due to a longer duration of device use. 396

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Valid** | **A – Well-** | **B – Meets or exceeds** | **C –** | **D – Associated** |
| **Predicate** | **established** | **expected predicate** | **Unmitigated** | **design-related** |
| **Device** | **methods** | **performance** | **use-related or** | **recall** |
|  |  |  | **design-related** |  |
|  |  |  | **safety issues** |  |
| 1 | Used relevant | Expected frequency of | No known | No design- |
|  | methods that | reported adverse | unmitigated use- | related recall |
|  | were published | events. History of safe | related or | identified |
|  | in the public | use established due to | design-related |  |
|  | domain | duration of device on | safety issues |  |
|  |  | the market. |  |  |
| 2 | Used relevant | Expected frequency of | No known | No design- |
|  | methods that | reported adverse events | unmitigated use- | related recall |
|  | were published |  | related or | identified |
|  | in the public |  | design-related |  |
|  | domain |  | safety issues |  |

397

1. In their draft 510(k) Summary, the submitter includes a brief narrative describing the above
2. selection process in the proposed 510(k) Summary. The submitter’s draft 510(k) Summary also
3. includes a discussion that the predicate used to support the 510(k) submission used well-
4. established methods for IFDs, discusses the frequency of reported adverse events, including that
5. the device has a well-established safety profile through a history of safe use due to its longer
6. duration on the market, and states that there are no known unmitigated use-related or design-
7. related safety issues or associated design-related recalls.